## Sheet1
Failure Modes Effects and Critical Analysis (FMECA)
Part: TetherIA Aero Hand Open |  |  | Version: 1st version |  | Lihao ZHOU |  | Date: 05/01/2026
Item / Component | Failure Mode | Failure Cause | Failure Effect | Severity (S) | Occurrence (O) | Detection (D) | RPN | Recommended Mitigation Action
Tendon Cables (Tension Lines) | Mechanical snapping or severe plastic deformation | High friction within routing channels; excessive pulling force from servos | Loss of actuation in the affected digit; severe mechanical hysteresis | 8 | 8 | 4 | 256 | Utilise high-strength UHMWPE lines; apply PTFE lubricant to internal routing paths.
Actuators (Servo Motors) | Gear stripping or thermal overload (burnout) | Continuous stall current when physical movement is blocked; internal friction | Permanent loss of motor function; potential electrical short circuit | 9 | 5 | 5 | 225 | Implement software-level angular constraints; upgrade to metal-gear servos.
Tendon Routing Channels | Tendon jamming / excessive dynamic friction | Rough internal surfaces from 3D printed support material; acute bending angles | Delayed finger response; increased electrical load and heat on servos | 6 | 9 | 3 | 162 | Conduct mechanical post-processing (sanding) of internal channels; minimise sharp routing angles.
Structural Fasteners (Pins/Screws) | Loosening or mechanical shear | Cyclic dynamic loading during continuous grasping cycles | Joint instability; imprecise kinematic output at the fingertips | 6 | 5 | 4 | 120 | Apply thread-locking fluid during assembly; utilise metal pins for primary rotary joints.
3D-Printed Phalanges (Finger Joints) | Structural fracture or layer delamination | Mechanical stress from tendon tension; poor interlayer adhesion during FDM printing | Complete loss of structural integrity; inability to perform flexion/extension | 8 | 6 | 2 | 96 | Increase 3D print infill density; optimise print orientation for load distribution.

## Sheet2
Rating | Severity (S) - Impact on System Operation | Occurrence (O) - Probability of Failure | Detection (D) - Likelihood of Detecting Before Failure
10 | Hazardous: Complete system failure; potential safety hazard. | Inevitable: Failure is almost certain to occur repeatedly. | Absolute Uncertainty: No method exists to detect the failure before it occurs.
9 | Critical: Primary function lost; requires complete hardware replacement. | Very High: Failures occur frequently during normal operation. | Very Low: Detection is highly improbable before catastrophic failure.
8 | High: Severe degradation of primary function; significant mechanical stall. | High: Repeated failures observed during testing phases. | Low: System has minimal capacity to detect the impending issue.
7 | Major: Partial loss of primary function; high execution hysteresis. | Moderately High: Failures occur occasionally. | Moderately Low: Detection requires specific, non-routine manual inspection.
6 | Significant: System operates but with noticeable physical degradation. | Moderate: Occasional failures, but statistically manageable. | Moderate: Issue can be detected during routine system calibration.
5 | Moderate: Minor loss of performance; necessitates hardware recalibration. | Occasional: Isolated incidents; not a persistent pattern. | Fair: Moderate probability of detection via software anomalies or physical visual checks.
4 | Minor: System functions with slight friction or delay; acceptable tolerance. | Low: Few failures expected during the component's lifespan. | Moderately High: High probability of detection during standard pre-flight checks.
3 | Slight: Negligible impact on overall data acquisition or kinematics. | Very Low: Failures are rare and highly unlikely. | High: Very likely to be detected immediately by system feedback.
2 | Very Slight: Barely noticeable effect; cosmetic or transient issue. | Remote: Extremely unlikely to occur under normal parameters. | Very High: Nearly certain to be detected by built-in continuous monitoring.
1 | None: No discernible impact on system functionality. | Improbable: History of no failures for similar engineering designs. | Certain: Automated and immediate detection guaranteed prior to operational failure.